## Summary
¦ | Scenario | ¦ | Description | ¦ | Good | ¦ | Decent | ¦ | Bad | ¦
¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦
¦ | Healthy PC-CMV | ¦ | PC-CMV mode | ¦<span class="success"> | 11 | </span>¦<span class="warning"> | 0 | </span>¦<span class="danger"> | 0 | </span>¦
¦ | Healthy VC-CMV | ¦ | VC-CMV mode | ¦<span class="success"> | 11 | </span>¦<span class="warning"> | 0 | </span>¦<span class="danger"> | 0 | </span>¦
¦ | Healthy VC-AC | ¦ | VC-AC mode | ¦<span class="success"> | 11 | </span>¦<span class="warning"> | 0 | </span>¦<span class="danger"> | 0 | </span>¦
¦ | Mild ARDS | ¦ | PC-CMV with Severity = 0.3 | ¦<span class="success"> | 11 | </span>¦<span class="warning"> | 0 | </span>¦<span class="danger"> | 0 | </span>¦
¦ | Moderate ARDS | ¦ | PC-CMV with Severity = 0.6 | ¦<span class="success"> | 11 | </span>¦<span class="warning"> | 0 | </span>¦<span class="danger"> | 0 | </span>¦
¦ | Severe ARDS | ¦ | PC-CMV with Severity = 0.9 | ¦<span class="success"> | 11 | </span>¦<span class="warning"> | 0 | </span>¦<span class="danger"> | 0 | </span>¦
¦ | Mild COPD | ¦ | PC-CMV with Severity = 0.3 | ¦<span class="success"> | 11 | </span>¦<span class="warning"> | 0 | </span>¦<span class="danger"> | 0 | </span>¦
¦ | Moderate COPD | ¦ | PC-CMV with Severity = 0.6 | ¦<span class="success"> | 11 | </span>¦<span class="warning"> | 0 | </span>¦<span class="danger"> | 0 | </span>¦
¦ | Severe COPD | ¦ | PC-CMV with Severity = 0.9 | ¦<span class="success"> | 11 | </span>¦<span class="warning"> | 0 | </span>¦<span class="danger"> | 0 | </span>¦
¦ | Recruitment | ¦ | Increase PEEP for improved recruitment | ¦<span class="success"> | 8 | </span>¦<span class="warning"> | 0 | </span>¦<span class="danger"> | 0 | </span>¦
¦ |  | ¦ | Total | ¦<span class="success"> | 107 | </span>¦<span class="warning"> | 0 | </span>¦<span class="danger"> | 0 | </span>¦

## Scenarios
¦ | Patient State | ¦ | Ventilator Settings | ¦ | Respiration Rate (bpm) | ¦ | Inspiratory-Expiratory Ratio | ¦ | Pulmonary Resistance (cmH2O-s/L) | ¦ | Pulmonary Compliance (L/cmH2O) | ¦ | pH | ¦ | Arterial Carbon Dioxide Pressure (mmHg) | ¦ | Arterial Oxygen Pressure (mmHg) | ¦ | PaO2/FiO2 (mmHg) | ¦ | Oxygen Saturation (%) | ¦ | Pulmonary Shunt Fraction (%) | ¦ | Alveolar Dead Space (L) | ¦
¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦
¦ | Healthy | ¦ | Mode: PC-CMV<br>Inspiratory Pressure Target: set empirically to achieve VT Target<br>VT Target (mL/kg): 7<br>VT Target (mL): 527<br>PEEP (cm H2O): 5<br>FiO2 (%): 30<br>Mandatory Rate (bpm): 15 <br>I:E: 0.60<br>Minute Ventilation (L/min): 7.9<br> | ¦<span class="success"> | 15 (ventilator setting) | </span>¦<span class="success"> | 0.6 (ventilator setting) | </span>¦<span class="success"> | 12-17 @cite arnal2013feasibility, @cite arnal2018parameters | </span>¦<span class="success"> | 20-54 @cite arnal2013feasibility, @cite arnal2018parameters | </span>¦<span class="success"> | 7.27-7.40 @cite arnal2013feasibility | </span>¦<span class="success"> | 35-42 @cite arnal2013feasibility | </span>¦<span class="success"> | 85-114 @cite arnal2013feasibility | </span>¦<span class="success"> | 264-366 @cite arnal2013feasibility, @cite ferguson2012berlin | </span>¦<span class="success"> | 96.4-98.5 @cite arnal2013feasibility, @cite kacmarek2016egan | </span>¦<span class="success"> | 2-5 @cite Levitzky2013pulmonary, @cite radermacher2017fifty, @cite petersson2014gas | </span>¦<span class="success"> | ~0 @cite Levitzky2013pulmonary | </span>¦
¦ | Healthy | ¦ | Mode: VC-CMV<br>Inspiratory Pressure Target: set empirically to achieve VT Target<br>VT Target (mL/kg): 7<br>VT Target (mL): 527<br>PEEP (cm H2O): 5<br>FiO2 (%): 30<br>Mandatory Rate (bpm): 15 <br>I:E: 0.60<br>Minute Ventilation (L/min): 7.9<br> | ¦<span class="success"> | 15 (ventilator setting) | </span>¦<span class="success"> | 0.6 (ventilator setting) | </span>¦<span class="success"> | 12-17 @cite arnal2013feasibility, @cite arnal2018parameters | </span>¦<span class="success"> | 20-54 @cite arnal2013feasibility, @cite arnal2018parameters | </span>¦<span class="success"> | 7.27-7.40 @cite arnal2013feasibility | </span>¦<span class="success"> | 35-42 @cite arnal2013feasibility | </span>¦<span class="success"> | 85-114 @cite arnal2013feasibility | </span>¦<span class="success"> | 264-366 @cite arnal2013feasibility, @cite ferguson2012berlin | </span>¦<span class="success"> | 96.4-98.5 @cite arnal2013feasibility, @cite kacmarek2016egan | </span>¦<span class="success"> | 2-5 @cite Levitzky2013pulmonary, @cite radermacher2017fifty, @cite petersson2014gas | </span>¦<span class="success"> | ~0 @cite Levitzky2013pulmonary | </span>¦
¦ | Healthy | ¦ | Mode: VC-AC<br>Inspiratory Pressure Target: set empirically to achieve VT Target<br>VT Target (mL/kg): 7<br>VT Target (mL): 527<br>PEEP (cm H2O): 5<br>FiO2 (%): 30<br>Mandatory Rate (bpm): 15 <br>I:E: 0.60<br>Minute Ventilation (L/min): 7.9<br> | ¦<span class="success"> | 15-21 @cite arnal2013feasibility | </span>¦<span class="success"> | 0.4-0.8 @cite arnal2013feasibility | </span>¦<span class="success"> | 12-17 @cite arnal2013feasibility, @cite arnal2018parameters | </span>¦<span class="success"> | 20-54 @cite arnal2013feasibility, @cite arnal2018parameters | </span>¦<span class="success"> | 7.27-7.40 @cite arnal2013feasibility | </span>¦<span class="success"> | 35-42 @cite arnal2013feasibility | </span>¦<span class="success"> | 85-114 @cite arnal2013feasibility | </span>¦<span class="success"> | 264-366 @cite arnal2013feasibility, @cite ferguson2012berlin | </span>¦<span class="success"> | 96.4-98.5 @cite arnal2013feasibility, @cite kacmarek2016egan | </span>¦<span class="success"> | 2-5 @cite Levitzky2013pulmonary, @cite radermacher2017fifty, @cite petersson2014gas | </span>¦<span class="success"> | ~0 @cite Levitzky2013pulmonary | </span>¦
¦ | Mild ARDS (severity = 0.3) | ¦ | Mode: PC-CMV<br>Inspiratory Pressure Target: set empirically to achieve VT Target<br>VT Target (mL/kg): 6<br>VT Target (mL): 452<br>PEEP (cm H2O): 11<br>FiO2 (%): 40<br>Mandatory Rate (bpm): 17<br>I:E: 0.90<br>Minute Ventilation (L/min): 7.6<br> | ¦<span class="success"> | 17 (ventilator setting) | </span>¦<span class="success"> | 0.9 (ventilator setting) | </span>¦<span class="success"> | No change, 10-16 @cite arnal2013feasibility, @cite arnal2018parameters | </span>¦<span class="success"> | Decrease, 20-54 @cite arnal2013feasibility, @cite arnal2018parameters | </span>¦<span class="success"> | 7.15-7.34 @cite arnal2013feasibility | </span>¦<span class="success"> | 45-57 @cite arnal2013feasibility | </span>¦<span class="success"> | 103-133 @cite arnal2013feasibility | </span>¦<span class="success"> | Decrease, 200-300 @cite ferguson2012berlin | </span>¦<span class="success"> | 98.0-99.0 @cite arnal2013feasibility, @cite kacmarek2016egan | </span>¦<span class="success"> | Increase, >10% @cite radermacher2017fifty, @cite petersson2014gas | </span>¦<span class="success"> | Increase @cite nuckton2002pulmonary | </span>¦
¦ | Moderate ARDS (severity = 0.6) | ¦ | Mode: PC-CMV<br>Inspiratory Pressure Target: set empirically to achieve VT Target<br>VT Target (mL/kg): 6<br>VT Target (mL): 452<br>PEEP (cm H2O): 11<br>FiO2 (%): 70<br>Mandatory Rate (bpm): 20<br>I:E: 0.90<br>Minute Ventilation (L/min): 9.0<br> | ¦<span class="success"> | 20 (ventilator setting) | </span>¦<span class="success"> | 0.9 (ventilator setting) | </span>¦<span class="success"> | No change, 10-16 @cite arnal2013feasibility, @cite arnal2018parameters | </span>¦<span class="success"> | Decrease, 20-54 @cite arnal2013feasibility, @cite arnal2018parameters | </span>¦<span class="success"> | 7.15-7.34 @cite arnal2013feasibility | </span>¦<span class="success"> | 45-57 @cite arnal2013feasibility | </span>¦<span class="success"> | 81-117 @cite arnal2013feasibility | </span>¦<span class="success"> | Decrease, 100-200 @cite ferguson2012berlin | </span>¦<span class="success"> | 95.9-98.6 @cite arnal2013feasibility, @cite kacmarek2016egan | </span>¦<span class="success"> | Increase, >20% @cite radermacher2017fifty, @cite petersson2014gas | </span>¦<span class="success"> | Increase @cite nuckton2002pulmonary | </span>¦
¦ | Severe ARDS (severity = 0.9) | ¦ | Mode: PC-CMV<br>Inspiratory Pressure Target: set empirically to achieve VT Target<br>VT Target (mL/kg): 6<br>VT Target (mL): 452<br>PEEP (cm H2O): 11<br>FiO2 (%): 100<br>Mandatory Rate (bpm): 25<br>I:E: 0.90<br>Minute Ventilation (L/min): 11.5<br> | ¦<span class="success"> | 25 (ventilator setting) | </span>¦<span class="success"> | 0.9 (ventilator setting) | </span>¦<span class="success"> | No change, 10-16 @cite arnal2013feasibility, @cite arnal2018parameters | </span>¦<span class="success"> | Decrease, 20-54 @cite arnal2013feasibility, @cite arnal2018parameters | </span>¦<span class="success"> | 7.15-7.34 @cite arnal2013feasibility | </span>¦<span class="success"> | 45-57 @cite arnal2013feasibility | </span>¦<span class="success"> | 56-83 @cite arnal2013feasibility | </span>¦<span class="success"> | Decrease, <100 @cite ferguson2012berlin | </span>¦<span class="success"> | 88.8-96.1 @cite arnal2013feasibility, @cite kacmarek2016egan | </span>¦<span class="success"> | Increase, >40% @cite radermacher2017fifty, @cite petersson2014gas | </span>¦<span class="success"> | Increase @cite nuckton2002pulmonary | </span>¦
¦ | Mild COPD (severity = 0.3) | ¦ | Mode: PC-CMV<br>Inspiratory Pressure Target: set empirically to achieve VT Target<br>VT Target (mL/kg): 9.0<br>VT Target (mL): 678<br>PEEP (cm H2O): 8<br>FiO2 (%): 31<br>Mandatory Rate (bpm): 12<br>I:E: 0.43<br>Minute Ventilation (L/min): 8.0<br> | ¦<span class="success"> | 12 (ventilator setting) | </span>¦<span class="success"> | 0.43 (ventilator setting) | </span>¦<span class="success"> | Increase, 17-26 @cite arnal2013feasibility, @cite arnal2018parameters | </span>¦<span class="success"> | Increase, 46-61 @cite arnal2013feasibility, @cite arnal2018parameters | </span>¦<span class="success"> | 7.2-7.31 @cite arnal2013feasibility | </span>¦<span class="success"> | 44-67 @cite arnal2013feasibility | </span>¦<span class="success"> | 77-94 @cite arnal2013feasibility | </span>¦<span class="success"> | Decrease, 206-295 @cite ferguson2012berlin | </span>¦<span class="success"> | 95.2-97.3 @cite arnal2013feasibility, @cite kacmarek2016egan | </span>¦<span class="success"> | Increase @cite boerrigter2011cardiac | </span>¦<span class="success"> | Increase @cite Levitzky2013pulmonary | </span>¦
¦ | Moderate COPD (severity = 0.6) | ¦ | Mode: PC-CMV<br>Inspiratory Pressure Target: set empirically to achieve VT Target<br>VT Target (mL/kg): 8.5<br>VT Target (mL): 640<br>PEEP (cm H2O): 8<br>FiO2 (%): 31<br>Mandatory Rate (bpm): 14<br>I:E: 0.43<br>Minute Ventilation (L/min): 9.0<br> | ¦<span class="success"> | 14 (ventilator setting) | </span>¦<span class="success"> | 0.43 (ventilator setting) | </span>¦<span class="success"> | Increase, 17-26 @cite arnal2013feasibility, @cite arnal2018parameters | </span>¦<span class="success"> | Increase, 46-61 @cite arnal2013feasibility, @cite arnal2018parameters | </span>¦<span class="success"> | 7.2-7.31 @cite arnal2013feasibility | </span>¦<span class="success"> | 44-67 @cite arnal2013feasibility | </span>¦<span class="success"> | 77-94 @cite arnal2013feasibility | </span>¦<span class="success"> | Decrease, 206-295 @cite ferguson2012berlin | </span>¦<span class="success"> | 95.2-97.3 @cite arnal2013feasibility, @cite kacmarek2016egan | </span>¦<span class="success"> | Increase @cite boerrigter2011cardiac | </span>¦<span class="success"> | Increase @cite Levitzky2013pulmonary | </span>¦
¦ | Severe COPD (severity = 0.9) | ¦ | Mode: PC-CMV<br>Inspiratory Pressure Target: set empirically to achieve VT Target<br>VT Target (mL/kg): 9.0<br>VT Target (mL): 678<br>PEEP (cm H2O): 8<br>FiO2 (%): 40<br>Mandatory Rate (bpm): 15<br>I:E: 0.43<br>Minute Ventilation (L/min): 10.2<br> | ¦<span class="success"> | 15 (ventilator setting) | </span>¦<span class="success"> | 0.43 (ventilator setting) | </span>¦<span class="success"> | Increase, 17-26 @cite arnal2013feasibility, @cite arnal2018parameters | </span>¦<span class="success"> | Increase, 46-61 @cite arnal2013feasibility, @cite arnal2018parameters | </span>¦<span class="success"> | 7.2-7.31 @cite arnal2013feasibility | </span>¦<span class="success"> | 44-67 @cite arnal2013feasibility | </span>¦<span class="success"> | 77-94 @cite arnal2013feasibility | </span>¦<span class="success"> | Decrease, 206-295 @cite ferguson2012berlin | </span>¦<span class="success"> | 95.2-97.3 @cite arnal2013feasibility, @cite kacmarek2016egan | </span>¦<span class="success"> | Increase @cite boerrigter2011cardiac | </span>¦<span class="success"> | Increase @cite Levitzky2013pulmonary | </span>¦

## Recruitment
¦ | Segment | ¦ | Notes | ¦ | Action Occurrence Time (s) | ¦ | Sampled Scenario Time (s) | ¦ | Pulmonary Shunt (%) | ¦ | Horowitz Index (mmHg) | ¦ | Arterial Oxygen Partial Pressure (mmHg) | ¦ | Arterial Carbon Dioxide Partial Pressure (mmHg) | ¦
¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦
¦ | Ventilate patient with moderate ARDS | ¦ | Apneic and intubated; VC-CMV: RR = 15 bpm, I:E = 0.6, TV = 7 mL/kg(ideal), FIO2 = 0.5, PEEP = 5.0 cmH2O | ¦ | 30 | ¦ | 330 | ¦<span class="success"> | 33 +/- 15 @cite karbing2020changes | ¦<span class="success"> | 130 +/- 58 @cite karbing2020changes | ¦<span class="success"> | Increase @cite karbing2020changes | ¦<span class="success"> | No Change @cite musch2008relation | </span>¦
¦ | Increase PEEP | ¦ | PEEP = 20.0 cmH2O | ¦ | 330 | ¦ | 630 | ¦<span class="success"> | 22 +/- 14 @cite karbing2020changes | </span>¦<span class="success"> | 220 +/- 82 @cite karbing2020changes | </span>¦<span class="success"> | Increase @cite karbing2020changes | </span>¦<span class="success"> | No Change @cite musch2008relation | </span>¦